Study of the D-D neutron generator in application to the ITER neutron diagnostics calibration [[1]](#footnote-1)\*)

DOI: 10.34854/ICPAF.2021.48.1.179

1,2Kormilitsyn T.M., 1Rodionov R.N., 1,2Pankratenko A.V., 1Obudovsky S.Yu., 1,2Kashchuk Yu.A.

1Project Center ITER (RF DA), Moscow, Russia, [T.Kormilitsyn@iterrf.ru](mailto:T.Kormilitsyn@iterrf.ru)  
2Moscow Institute of Physics and Technology (National Research University), Dolgoprudny,  
 Russia

Current research focuses on the D-D neutron generator study in the view of utilization of these sources during the ITER neutron diagnostics *in situ* calibration campaigns. The results of anisotropy analysis in terms of both yield and energy distribution are presented using the ING-07D neutron generator as an example.

We conducted the GEANT4-modelling [1] of the LaCl3(Ce) scintillation detector response to the neutron source. Parameters of the neutron source were selected in accordance with the realistic ING-07D operating regime. We performed the comparative analysis of the experimentally obtained detector responses and the modelled ones [2]. The digital pulse-shape discrimination method was applied allowing the in-depth analysis of only the neutron-induced LaCl3(Ce) response part [3].

Results of this analysis allowed us to infer the ING-07D ion beam composition. Specific features of the neutron energy distribution are to be considered during the ITER in-vessel neutron calibration campaign employing this type of neutron generators.

Based on the experimental data we were able to estimate the D-D neutron generator directional properties in terms of yield, several operational parameters to be controlled during calibration campaigns were identified. Additional study was conducted featuring the long-term stability of the neutron generator yield.

The D-D neutron generator type in scope of current research is shown to be the optimal solution to the task of the ITER in-vessel neutron diagnostics calibration. The LaCl3(Ce) scintillation detector is deemed suitable for the purposes of D-D neutron generator yield and spectrum monitoring.

Current research is funded under the State Contract №Н.4а.241.19.20.1042 as of 21 April 2020.

References

1. J. Allison et al., “Recent developments in GEANT4”, *Nucl. Instruments Methods Phys. Res. Sect. A Accel. Spectrometers, Detect. Assoc. Equip.*, vol. 835, pp. 186–225, Nov. 2016.
2. T.M. Kormilitsyn et al., “On the application of the LaCl3(Ce) scintillation detector to the fast neutron detection”, *Physics of Elementary Particles and Atomic Nuclei, Letters*, 2021, Т.18, №1(233), С 86-97.
3. Pankratenko A.V. et al., «Analysis of the LaCl3 characteristics for the purpose of deuterium plasma neutron diagnostics», XLVIII Zvenigorod Conference on the Plasma Physics and the Controlled Fusion, 15 – 19 March, 2021.

1. \*) [abstracts of this report in Russian](http://www.fpl.gpi.ru/Zvenigorod/XLVIII/E/ru/HU-Kormilitsin.docx) [↑](#footnote-ref-1)